## Consolidated Detailed Budget
GUYANA
NATIONAL QUALITY INFRASTRUCTURE
(GY-ll1059)
CONSOLIDATED BUDGET
(US Dollars)
Description/Source | IDB
I. DIRECT COSTS | 7569500
A. sub-Component 1: Modernization of the Institutional Framework of the NQI | 615000
Product 1.1. Accreditation of existing laboratories | 200000
1.1.1. Specialized Consultancies for Accreditation | 200000
Product 1.2. Training and Capacity Building for Stakeholders | 105000
1.2.1. Short Training Courses | 30000
1.2.2. Specialized Training Courses | 75000
Product 1.3. Creation and Operation of a Laboratory Network | 135000
1.3.1. Awarennes and Sensitivation Worshops (3) | 20000
1.3.2. Specialized Consultancy for Information and Analysis | 50000
1.3.4. Design and Implementation of Website | 25000
1.3.5. Operation of the Network | 40000
Product 1.4. Capacity Building for SMES and private sector | 75000
1.4.1. Short training seminars (5) | 50000
1.4.2. Awarenness Campaings (1) | 25000
Product 1.5. Strenghen the Capacity of the GNBS | 100000
1.5..1. Specialized Consultancy for the Strategic Plan | 50000
1.5.2. Specialized Training (4) | 50000
B. Sub- Component 2: Improving Laboratory Facilities and Equipment | 5794500
Product 2.1. Arquitectural and Engineering Designs | 150000
2.1.2. Specialized Consultancy for architectural and engineering designs | 150000
Product 2.2. Facility for the GNBS | 4050000
2.2.1. Civil Works | 3400000
2.2.2. Civil Works Supervision and Inspection | 200000
2.2.3. Social and Environmental Supervision | 150000
2.2.4. Laboratory Furniture | 200000
2.2.5. Office Equipment | 100000
Product 2.3. GNBS Equipment | 1504500
2.3.1. Air Conditioning System | 300000
2.3.2. Fire System | 60000
2.3.3. Specialized Equipment and Patterns (*) | 594500
2.3.4. Electronic System (Includes Generator) | 400000
2.3.5. Solar Panels | 100000
2.3.5. Maintenance Contracts | 50000
Product 2.4. Specifc Gaps in other Laboratories | 90000
2.4. Specialized Equipment | 90000
C. Sub-Component 3. Implement National Export and Investment Strategy | 1160000
Product 3.1. Strenghening Go-Invest Trade and Investment Framework | 760000
3.1.1. Best Practices on IPAS | 50000
3.1.2. Redesign of Go-Invest and Trade Investement Framework | 510000
3.1.3. Customer Relationship Management and Training | 40000
3.1.4. Website Improvement | 70000
3.1.5. Boutique Forum | 90000
Product 3.2. Export and Investment Strategy | 400000
II. PROGRAM ADMINISTRATION | 1280000
A. Personnel | 1230000
1. Management and Coordination | 708000
2. Finance and Admin | 312000
3. Technical Support | 210000
B. Equipment, Logistics and Space | 50000
1. Office Furniture | 50000
IV. MONITORING AND EVALUATION | 50000
A. Midterm and final evaluation | 20000
B. Base line Collection | 15000
C. ex post cost benefit analysis | 15000
V. CONTINGENCIES | 100500
T O T A L | 9000000

## AOP (2)
Description/Source | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Year 6 | IDB | Number | Unit Cost
Year 6
I. DIRECT COSTS | 774000 | 2261000 | 3173900 | 846700 | 493900 | 20000 | 7569500 | 38 | 7094500 | 7569500
A. sub-Component 1: Modernization of the Institutional Framework of the NQI | 210000 | 155000 | 155000 | 60000 | 30000 | 5000 | 615000 | 21 | 335000 | 615000
Product 1.1. Accreditation of existing laboratories | 50000 | 50000 | 50000 | 25000 | 25000 | 0 | 200000 | 4 | 50000 | 200000
1.1.1. Specialized Consultancies for Accreditation | 50000 | 50000 | 50000 | 25000 | 25000 | 0 | 200000 | 4 | 50000 | 200000
Product 1.2. Training and Capacity Building for Stakeholders | 10000 | 35000 | 35000 | 25000 | 0 | 0 | 105000 | 6 | 35000 | 105000
1.2.1. Short Training Courses | 10000 | 10000 | 10000 |  |  |  | 30000 | 3 | 10000 | 30000
1.2.2. Specialized Training Courses |  | 25000 | 25000 | 25000 |  |  | 75000 | 3 | 25000 | 75000
Product 1.3. Creation and Operation of a Laboratory Network | 75000 | 20000 | 20000 | 10000 | 5000 | 5000 | 135000 | 5 | 125000 | 135000
1.3.1. Awarennes and Sensitivation Worshops (3) |  | 10000 | 10000 |  |  |  | 20000 | 2 | 10000 | 20000
1.3.2. Specialized Consultancy for Information and Analysis | 50000 | [5 empty] | 50000 | 1 | 50000 | 50000
1.3.3. Design and Implementation of Website | 25000 | [5 empty] | 25000 | 1 | 25000 | 25000
1.3.4. Operation of the Network |  | 10000 | 10000 | 10000 | 5000 | 5000 | 40000 | 1 | 40000 | 40000
Product 1.4. Capacity Building for SMES and private sector | 25000 | 25000 | 25000 | 0 | 0 | 0 | 75000 | 3 | 50000 | 75000
1.4.1. Short training seminars | 25000 | 25000 |  |  |  |  | 50000 | 2 | 25000 | 50000
1.4.2. Awarenness Campaings |  |  | 25000 |  |  |  | 25000 | 1 | 25000 | 25000
Product 1.5. Strenghen the Capacity of the GNBS | 50000 | 25000 | 25000 | 0 | 0 | 0 | 100000 | 3 | 75000 | 100000
1.5.1. Specialized Consultancy for the Strategic Plan | 50000 | [5 empty] | 50000 | 1 | 50000 | 50000
1.5.2. Specialized Training |  | 25000 | 25000 |  |  |  | 50000 | 2 | 25000 | 50000
B. Sub- Component 2: Improving Laboratory Facilities and Equipment | 150000 | 1470000 | 2908900 | 786700 | 463900 | 15000 | 5794500 | 13 | 5599500 | 5794500
Product 2.1. Arquitectural and Engineering Designs | 150000 | 0 | 0 | 0 | 0 | 0 | 150000 | 1 | 150000 | 150000
2.1.2. Specialized Consultancy for architectural and engineering designs | 150000 | [5 empty] | 150000 | 1 | 150000 | 150000
Product 2.2. Facility for the GNBS | 0 | 1470000 | 2310000 | 150000 | 105000 | 15000 | 4050000 | 4 | 3900000 | 4050000
2.2.1. Civil Works |  | 1360000 | 2040000 |  |  |  | 3400000 | 1 | 3400000 | 3400000
2.2.2. Civil Works Supervision and Inspection |  | 80000 | 120000 |  |  |  | 200000 | 1 | 200000 | 200000
2.2.3. Social and Environmental Supervision |  | 30000 | 50000 | 40000 | 15000 | 15000 | 150000 | 1 | 150000 | 150000
2.2.4. Laboratory Furniture |  |  | 50000 | 60000 | 90000 | 0 | 200000 | 1 | 200000 | 200000
2.2.5. Office Equipment |  |  | 50000 | 50000 |  |  | 100000 | 1 | 100000 | 100000
Product 2.3. GNBS Equipment | 0 | 0 | 598900 | 636700 | 268900 | 0 | 1504500 | 6 | 1504500 | 1504500
2.3.1. Air Conditioning System |  |  | 300000 |  |  |  | 300000 | 1 | 300000 | 300000
2.3.2. Fire System |  |  | 30000 | 30000 |  |  | 60000 | 1 | 60000 | 60000
2.3.3. Specialized Equipment and Patterns (*) |  |  | 118900 | 356700 | 118900 | 0 | 594500 | 1 | 594500 | 594500
2.3.4. Electronic System (Includes Generator) |  |  | 100000 | 200000 | 100000 | 0 | 400000 | 1 | 400000 | 400000
2.3.5. Solar Panels |  |  | 50000 | 50000 |  |  | 100000 | 1 | 100000 | 100000
2.3.6. Maintenance Contracts |  |  |  |  | 50000 | 0 | 50000 | 1 | 50000 | 50000
Product 2.4. Specifc Gaps in other Laboratories |  |  |  |  | 90000 | 0 | 90000 | 2 | 45000 | 90000
2.4. Specialized Equipment |  |  |  |  | 90000 | 0 | 90000 | 2 | 45000 | 90000
C. Sub-Component 3. Implement National Export and Investment Strategy | 414000 | 636000 | 110000 | 0 | 0 | 0 | 1160000 | 4 | 1160000 | 1160000
Product 3.1. Strenghening Go-Invest Trade and Investment Framework | 254000 | 396000 | 110000 | 0 | 0 | 0 | 760000 | 3 | 760000 | 760000
3.1.1. Best Practices on IPAS | 50000 | [5 empty] | 50000 | 1 | 50000 | 50000
3.1.2. Redesign of Go-Invest and Trade Investement Framework | 204000 | 306000 |  |  |  |  | 510000 | 1 | 510000 | 510000
3.1.3. Customer Relationship Management and Training |  | 20000 | 20000 |  |  |  | 40000 | 1 | 40000 | 40000
3.1.4. Website Improvement |  | 70000 |  |  |  |  | 70000 | 1 | 70000 | 70000
3.1.5. Boutique Forum |  |  | 90000 |  |  |  | 90000 | 1 | 90000 | 90000
Product 3.2. Export and Investment Strategy | 160000 | 240000 | 0 | 0 | 0 | 0 | 400000 | 1 | 400000 | 400000
II. PROGRAM ADMINISTRATION | 173000 | 246000 | 246000 | 246000 | 246000 | 123000 | 1280000 | 3 | 1280000 | 1280000
A. Personnel | 123000 | 246000 | 246000 | 246000 | 246000 | 123000 | 1230000 | 2 | 1230000 | 1230000
1. Management and Coordination | 70800 | 141600 | 141600 | 141600 | 141600 | 70800 | 708000 | 0 | 708000 | 708000
2. Finance and Admin | 31200 | 62400 | 62400 | 62400 | 62400 | 31200 | 312000 | 1 | 312000 | 312000
3. Technical Support | 21000 | 42000 | 42000 | 42000 | 42000 | 21000 | 210000 | 1 | 210000 | 210000
B. Equipment, Logistics and Space | 50000 | 0 | 0 | 0 | 0 | 0 | 50000 | 1 | 50000 | 50000
1. Office Furniture | 50000 | [5 empty] | 50000 | 1 | 50000 | 50000
III. MONITORING AND EVALUATION | 15000 | 0 | 10000 | 0 | 0 | 25000 | 50000 | 4 | 40000 | 50000
A. Midterm and final evaluation |  |  | 10000 |  | 0 | 10000 | 20000 | 2 | 10000 | 20000
B. Base line Collection | 15000 | [5 empty] | 15000 | 1 | 15000 | 15000
C. ex post cost benefit analysis |  |  |  |  | 0 | 15000 | 15000 | 1 | 15000 | 15000
IV. CONTINGENCIES |  |  |  | 50250 | 50250 | 0 | 100500 | 1 | 100500 | 100500
T O T A L | 962000 | 2507000 | 3429900 | 1142950 | 790150 | 168000 | 9000000 | 46 | 8515000 | 9000000
(*) See detailed specialized equipment
DISBURSEMENTS
Description/Source | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Year 6 | Total
Year 6
IDB | 962000 | 2507000 | 3429900 | 1142950 | 790150 | 168000 | 9000000
% | 10.69% | 27.86% | 38.11% | 12.7% | 8.779% | 1.867% | 100%
Total | 962000 | 2507000 | 3429900 | 1142950 | 790150 | 168000 | 9000000

## AOP Physical
Description/Source | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | Year 6 | Number | Unit Cost
I. DIRECT COSTS | 8 | 11 | 13 | 5 | 8 | 1 | 46 | #REF! | 46
A. sub-Component 1: Modernization of the Institutional Framework of the NQI | 5 | 5 | 6 | 2 | 3 | 0 | 21 | 335000 | 21
Product 1.1. Accreditation of existing laboratories | 0 | 1 | 1 | 1 | 1 | 0 | 4 | 50000 | 4
1.1.1. Specialized Consultancies for Accreditation | 0 | 1 | 1 | 1 | 1 | 0 | 4 | 50000 | 4
Product 1.2. Training and Capacity Building for Stakeholders | 1 | 1 | 2 | 1 | 1 | 0 | 6 | 35000 | 6
1.2.1. Short Training Courses | 1 | 1 | 1 | 0 | 0 | 0 | 3 | 10000 | 3
1.2.2. Specialized Training Courses | 0 | 0 | 1 | 1 | 1 | 0 | 3 | 25000 | 3
Product 1.3. Creation and Operation of a Laboratory Network | 2 | 1 | 1 | 0 | 1 | 0 | 5 | 125000 | 5
1.3.1. Awarennes and Sensitivation Worshops | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 10000 | 2
1.3.2. Specialized Consultancy for Information and Analysis | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 50000 | 1
1.3.3. Design and Implementation of Website | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 25000 | 1
1.3.4. Operation of the Network | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 40000 | 1
Product 1.4. Capacity Building for SMES and private sector | 1 | 1 | 1 | 0 | 0 | 0 | 3 | 50000 | 3
1.4.1. Short training seminars | 1 | 1 | 0 | 0 | 0 | 0 | 2 | 25000 | 2
1.4.2. Awarenness Campaings | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 25000 | 1
Product 1.5. Strenghen the Capacity of the GNBS | 1 | 1 | 1 | 0 | 0 | 0 | 3 | 75000 | 3
1.5.1. Specialized Consultancy for the Strategic Plan | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 50000 | 1
1.5.2. Specialized Training | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 25000 | 2
B. Sub- Component 2: Improving Laboratory Facilities and Equipment | 1 | 0 | 3 | 3 | 5 | 1 | 13 | 5749500 | 13
Product 2.1. Arquitectural and Engineering Designs | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 200000 | 1
2.1.2. Specialized Consultancy for architectural and engineering designs | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 200000 | 1
Product 2.2. Facility for the GNBS | 0 | 0 | 2 | 1 | 1 | 0 | 4 | 4000000 | 4
2.2.1. Civil Works | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 3500000 | 1
2.2.2. Civil Works Supervision and Inspection | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 200000 | 1
2.2.3. Social and Environmental Supervision | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 150000 | 1
2.2.4. Laboratory Furniture | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 200000 | 1
2.2.5. Office Equipment | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 100000 | 1
Product 2.3. GNBS Equipment | 0 | 0 | 1 | 2 | 2 | 1 | 6 | 1504500 | 6
2.3.1. Air Conditioning System | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 300000 | 1
2.3.2. Fire System | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 60000 | 1
2.3.3. Specialized Equipment and Patterns (*) | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 594500 | 1
2.3.4. Electronic System (Includes Generator) | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 400000 | 1
2.3.5. Solar Panels System |  |  | 0 | 1 | 0 | 0 | 1 | 100000 | 1
2.3.6. Maintenance Contracts |  |  |  |  | 0 | 1 | 1 | 50000 | 1
Product 2.4. Specifc Gaps in other Laboratories |  |  |  |  | 2 | 0 | 2 | 45000 | 2
2.4. Specialized Equipment |  |  |  |  | 2 | 0 | 2 | 45000 | 2
C. Sub-Component 3. Implement National Export and Investment Strategy | 2 | 6 | 4 | 0 | 0 | 0 | 12 | #REF! | 12
Product 3.1. Strenghening Go-Invest Trade and Investment Framework | 1 | 3 | 2 | 0 | 0 | 0 | 6 | 760000 | 6
3.1.1. Best Practices on IPAS | 1 | [5 empty] | 1 | 50000 | 1
3.1.2. Redesign of Go-Invest and Trade Investement Framework | 0 | 1 |  |  |  |  | 1 | 510000 | 1
3.1.3. Customer Relationship Management and Training |  | 0 | 1 |  |  |  | 1 | 40000 | 1
3.1.4. Website Improvement |  | 1 |  |  |  |  | 1 | 70000 | 1
3.1.5. Boutique Forum |  |  | 1 |  |  |  | 1 | 90000 | 1
Product 3.2. Export and Investment Strategy | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 400000 | 1
(1) See detailed specialized equipment

## PRESUP.CONS. (2)
GUYANA
NATIONAL QUALITY INFRASTRUCTURE
(GY-ll1059)
CONSOLIDATED BUDGET
(US Dollars)
Description/Source | IDB | Percentage (IDB)
I. DIRECT COSTS | 7569500 | 84.11% | 7569500 | 84.11%
A. Sub-Component 1: Modernization of the Institutional Framework of the NQI | 615000 | 6.833%
B. sub- Component 2: Improving Laboratory Facilities and Equipment | 5794500 | 64.38%
C. Sub-Component 3. Implement National Export and Investment Strategy | 1160000 | 12.89%
II. PROGRAM ADMINISTRATION | 1280000 | 14.22% | 1280000 | 14.22%
A. Personnel | 1230000 | 13.67%
B. Equipment, Logistics and Space | 50000 | 0.5556%
IV. MONITORING AND EVALUATION | 50000 | 0.5556%
V. CONTINGENCIES | 100500 | 1.117%
T O T A L | 9000000 | 100% | 9000000 | 100%
Operative costs budget as cost for provision of utilities and communications, office space and logistics and its maintenance.

## AdMin Costs
INTERAMERICAN DEVELOPMENT BANK
Guyana
(GY-L1059)
ADMINISTRATIVE BUDGET - PERSONNEL
(US Dollars)
Position | Monthly | Annual | Contract time | Number of | Total | Comments | Exchange Rate
| Salary | Salary |  | Months |  |  | 0.0048
I. COORDINATION AND MANAGEMENT |  |  |  |  | 708000
A. Program Coordinator | 4200 | 50400 | Whole Program | 60 | 252000
B. Technical Coordinator | 3000 | 36000 | Whole Program (2) | 120 | 360000
C. Monitoring and Evaluation Specialist | 1600 | 19200 | Part time | 60 | 96000
II. ADMINISTRACION Y FINANZAS | 3500 |  |  |  | 312000
A. Fiduciary Specialist | 2500 | 30000 | Whole Program | 60 | 150000
B. Procurement Specialist | 1800 | 21600 | Whole Program | 60 | 108000
D. Administrative Official | 900 | 10800 | Whole Program | 60 | 54000 | Part time
II. Technical Support |  |  |  |  | 210000
A. Technical Support/Engineering | 3000 | 36000 | Initial Program | 30 | 90000 | Supervision of works
A. Technical Support/Equipment | 3000 | 36000 | Initial Program | 20 | 60000 | Support on Specications
A. Technical Support/Environmental | 3000 | 36000 | Initial Program/part time | 20 | 60000 | Supervision of works
T O T A L |  |  |  |  | 1230000
1215000

## Equipment (2)
Item | Costo | Numero de | Total | Notas/
| Unitario | Unidades |  | Supuestos
I. Office Equipment for GNBS |  |  | 31000
A. Laptops | 2000 | 10 | 20000
B. Laser color printer | 1000 | 2 | 2000
C. Projector | 1500 | 1 | 1500
D. Various additional equipment | 7500 | 1 | 7500
II. SOFTWARE |  |  | 19000
A. Microsoft Office Suite 365 | 150 | 20 | 3000 |  | EACH LICENCE IS 150/YEAR/person, CONSIDERING LICENCES FOR 3 YEARS= 7 STAFF X 3 YEARS, Source http://office.microsoft.com/en-us/business/compare-all-office-365-for-business-plans-FX104051403.aspx
B. Stata Statistical Software (Stata 13) | 1000 | 6 | 6000 |  | ESTIMATED, MAY CHANGE BASED ON THE TYPE OF LICENSE
C. Various--business and statistical | 10000 | 1 | 10000
III Laboratory FURNITURE |  |  | 100000
A. Laboratory Furniture |  |  | 100000
I V. OFFICE INFRASTRUCTURE |  |  | 50000 | 200000
A. Office infrastructure | 50000 | 1 | 50000 |  | CUBICLES, FLOORS, OFFICE SPACE
0
V. Laboratory Equpipment |  |  | 1544500
A. Air conditioning System | 200000 | 1 | 200000
B. Fire System | 60000 | 1 | 60000
C. Electric System | 100000 | 4 | 400000
D. Specialized Laboratory Equipment | 594500 | 1 | 594500
Small Mass Laboratory | 20400 | 1 | 20400
Large Mass Laboratory | 6700 | 1 | 6700
Humidty and Temperature | 319500 | 1 | 319500
Volume | 92000 | 1 | 92000
Time and Frequency | 15200 | 1 | 15200
Electrical Magnitude | 24000 | 1 | 24000
Length | 14000 | 1 | 14000
Strengh and Torque | 58000 | 1 | 58000
Physics and Chemistry | 44700 | 1 | 44700
E. Solar Panels | 200000 | 1 | 200000
F. Specialized Equipment | 45000 | 2 | 90000
VI. Hardware Upgrade for Go-Invest |  |  | 55000
A. Computers | 2000 | 20 | 40000
B. Printers | 1000 | 3 | 3000
C. Projectors | 1500 | 3 | 4500
D. Additional Equipment | 7500 | 1 | 7500
TOTAL |  |  | 1744500

## Specialized Laboratory Equi (2
No. | Item | Specifications | Qty. | Total US$ | Comments:
1 | Table for mas comparators | Table designed for ensuring stable operating conditions for high precision measurements. The table top must be manufactured from polished granite and it is based on mild steel, metal construction. Robust construction of the table along with special vibration absorbers must ensure excellent isolation of measuring instruments from possible vibrations, sourcing from environmental factors. | 3 | 15
2 | Laboratory environmental condition sensors | Temperature: +5 °C to +45°C d=0,01°C accuracy:0,1°C Pressure: 850 hPa to 1050 hPa d=0,1 hPa accuracy: 2hPa | 1 | 2
Humidity. 20% to 100%d=0,1% accuracy: 2%
With data logger
RS232 connection
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
3 | Gaussmeter (Or similar device) | Portable gaussmeter (with hall probe or Fluxgate magnetometer) with indication and Min/Max function | 1 | 1.2
· Range: (25 – 3000) mT
· Accuracy: ≤0.05% of reading
The gaussmeter (or teslameter) shall be suitable for measuring permanent magnetization of weights ³1g, Class F1 to M3 according to OIML R111-1.
A valid calibration certificates from an accredited calibration laboratory or National Metrology Institute shall be provided for the gaussmeter.
4 | Permeability Indicator (Or similar device) | Permeability indicating device basing on the attracting method (OIML-R111-2004, B.6.5), that determines the relative magnetic permeability. | 1 | 0.7
Range: 13 standards (screw inserts) of rel. magnetic permeability from 1.01 to 5.0
The device shall be suitable for measuring permeability/susceptibility of weights ≥ 20 g of low permeability material according to OIML R111-1.
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
5 | Laptop Computer | Minimum Screen Size: 15 inches | 1 | 1.5
Minimum Processor Speed: 2.6 GhZ
Minimum RAM: 6GB
Minimum Hard Drive: 500 GB
Optical DVD Drive
English Keyboard
Operating System: Windows 7 or higher
20.4
Lista de equipos requeridos para el Laboratorio de Masas grandes
1 | Laboratory environmental condition sensors | Temperature: +5 °C to +45°C d=0,01°C accuracy:0,1°C Pressure: 850 hPa to 1050 hPa d=0,1 hPa accuracy: 2hPa | 1 | 2
Humidity. 20% to 100%d=0,1% accuracy: 2%
With data logger
RS232 connection
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
2 | Laptop Computer | Minimum Screen Size: 15 inches | 1 | 1.5
Minimum Processor Speed: 2.6 GhZ
Minimum RAM: 6GB
Minimum Hard Drive: 500 GB
Optical DVD Drive
English Keyboard
Operating System: Windows 7 or higher
2 | Vertically motorized crane | Capacity: 1200 kg | 1 | 3.2
6.7
Lista de equipos requeridos para el Laboratorio de Temperatura y humedad
No. | Item | Specifications | Qty. | Budget | Comments:
1 | Calibration Thermostatic Bath | Temp Range: -40°C to 125°C | 1 | 18 | See model 785M by ISOTECH or its equivalent by Fluke.
Parallel tube liquid
Volume: 98mm diameter, 485mm deep
Stability: ±0.006°C @ -40°C (methanol)
Fluids: methanol/ water
Vertical uniformity: <1 mK
Display resolution: 0,01 °C
Accessories:
Liquid in Glass Thermometer Support Kit.
SPRT and thermocouple Support kit.
Monocular and Support. Useful for viewing and magnifying the liquid column within the thermometer being calibrated.
10 L of methanol for reposition
2 | Calibration Thermostatic Bath | Temp Range: 30°C to 300°C | 1 | 18 | See model 785H by ISOTECH or its equivalent by Fluke
Parallel tube liquid
Volume: 98mm diameter, 485mm deep
Stability: ±0.006°C @ 125°C (oil)
Fluids: water/ oil
Vertical uniformity: <1 mK
Display resolution: 0,01 °C
10 L of oil for reposition
3 | Fluidized calibration bath | Temperature Range: 50°C to 700°C | 1 | 25 | See model:875 by ISOTECH
Volume : 67mm diameter, 475mm deep
Absolute Stability:±0.005°C at 150°C
In Equalizing Block: ±0.005°C at 450°C and ±0.020°C at 660°C in 10 minutes.
Vertical Uniformity: ±0.005°C at 350°C, ±0.005°C at 420°C, ±0.026°C at 660°C.
Compressed Air Supply: 1 BAR, 30 L/min at 100°C (50L max)
Accessories:
Adjustable height aluminium bronze equalising block.
Eight 8mm pockets x 240mm deep.
Oil free compressor and connecting tube. Used to supply air to the fluidized bed.
Liquid in Glass Thermometer Support Gantry System consisting of a multi-tube probe holder with ten 10.80mm by 470mm deep pockets, four support pillars and an adjustable stainless steel gantry with holes corresponding to the probe holder, to locate up to ten thermometers simultaneously.
4 | Dry well reference Ice point | Temperature range 0 °C (32 °F) | 1 | 5 | See model: 9101 by Fluke or equivalent.
Stability ±0.005 °C
Total instrument error ±0.02 °C, typical; ±0.05 °C max. (18–25 °C ambient)
Temperature Coefficient ±0.005 °C/ °C
Valid calibration certificates from an accredited calibration laboratory or National Metrology Institute shall be provided.
5 | Standard thermocouple type S | Hot Sheath Temperature Range: 600 °C to 1200°C | 2 | 6 | See Isotech or equivalent
The measuring assembly comprises a 7mm Dia. x 300mm length connected to the reference sheath which is a 4.5mm x 250mm stainless steel sheath suitable for referencing in a 0°C reference system
Valid calibration certificates from an accredited calibration laboratory1 or National Metrology Institute shall be provided.
6 | Three heater zone | Furnace temperature range 300 °C to 1200 °C | 1 | 14 | See model: 9118A-ITB by Fluke or equivalent.
| Furnace | Radial uniformity:
|  | 300 °C ± 0.1 °C |  |  | For thermocouple calibration.
700 °C ± 0.20 °C
1200 °C ± 0.25 °C
Axial uniformity at Full range: ± 0.2 °C
Temperature stability: ± 0.1 °C
Accessories: Isothermal ceramic block with 4 holes.
7 | Precision thermometer bridge | Probes Supported PRT’s, Thermistors & Thermocouples | 1 | 16 | See: micro K model 250 by ISOTECH or equivalent, Super thermometer model 1594A by fluke is not same.
Accuracy Whole Range (SPRT Ro ≥ 2.5Ω):0.250 ppm
|  | Equivalent Temperature Accuracy: 0.06mK |  |  | Notice micro K can read SPRT and thermocouples as well.
Resolution: 0.01mK
Resolution Voltage: 10nV
Accuracy - Voltage uncertainty:
Resistance Range 0-500 kΩ
Accepted external RS range 1Ω-10 kΩ,
Voltage Range (Thermocouple): ±125mV
Internal Resistance Standards: 1,10,25, 100, 400Ω
Interfaces:RS232, GPIB, USB
3 input Channels
Internal Data: 2Gb
Units: Ratio, V, Ω, °C, °F, K
Display mode: numerical and graphic
8 | Standard Platinum Resistance Thermometer (SPRT 25) | RTPW 25.5Ω | 2 | 12 | See model 909 H/25 ohm by ISOTECH or equivalent.
|  | Nominal Resistance 25.5Ω at 0°C |  |  | These two SPRT will be used as the reference standards in the lab.
Max. Current mA 1
Nominal Sensitivity 0.1 Ω/°C
Resistance Ratio Wga >1.11807
Self-heating 1mK / 25 microwatts
4 wire
Metal sheathed
Temperature range: -45°C to 670°C
Valid calibration certificates from an accredited calibration laboratory1 or National Metrology Institute shall be provided.
9 | Semi standard platinum resistance thermometer | Temperature range: -45°C to 450°C | 4 | 12 | Working standard to be used in field calibration.
|  | Diameter: 6mm |  |  | See: ISOTECH 935-14-16 or equivalent.
Length: 350mm
|  | Sensing length: 25mm |  |  | Two of them to be used as working standard in the lab.
Ro 100Ω ± 0.05 Ω
Alpha 0.003850 ± 0.000005
Standard IEC 60751
Stability: 0.010 Ω/year
Recommended Current: 1mA
Self-Heating at 1mA 0.004°C
Connection: Four Wire
Valid calibration certificates from an accredited calibration laboratory or National Metrology Institute shall be provided.
10 | Wireless Temperature/ Humidity monitoring system | Temperature range: -45 to 140°C | 10 | 22 | See: Val probe or equivalent
Temperature sensor: RTD
|  | Temperature accuracy: 0,1 °C |  |  | To be used as working standard for climatic chambers characterization.
Relative Humidity range: 20 to 100%
Humidity accuracy: 2%
316 L stainless steel body
Reader station with 10 positions.
Data acquisition software
11 | Precision thermometer (indicator) | 2 Channels PRTs, IEC60751 (2008), Callendar-Van Dusen, ITS90 | 1 | 8 | See: mili K by Isotech or
|  | 1mA and 1.428mA |  |  | Chub-E4 by Fluke or equivalent.
|  | 2 Channels for Thermocouples |  |  | To be used In field with working standard for characterization of: Thermostatic baths, furnaces and dry blocks.
Thermocouples:IEC584-1 1995 (B, E,J, K,N, R,S,T)
Resolution:
Resistance (PRTs): 0.00001Ω
(Thermistors): 0.001Ω
(Voltage): 0.00001mV
(Current): 0.001mA
(Temperature): 0.0001°C
Logging Capacity to store > 180 Days of time
stamped measurements to internal
memory
Input Connectors:
SPRTs/PRTs: LemoEPG.1B.306.
HLN 6-pin gold plated contacts
Thermocouples: Miniature thermocouple socket (ASTM E 1684-05)
4-20mA: 4mm sockets
12 | Two pressure humidity generator | Relative humidity range: 10% to 95% | 1 | 140 | See: Thunder Scientific, model 2500.
Temperature range: 0°C to 70 °C
Relative humidity uncertainty: 0,5%
Temperature Uncertainty: 0,06°C
Temperature uniformity: 0,1 °C
Display resolution: 0.01 °C
Chamber window: 152mm x 152mm
13 | Chilled Mirror | Frost/Dew Point range: -20…70 °C | 1 | 20 | See: RH System, model: 473.
Relative humidity range: 2…100 %
Temperature (head mount PRT): -20…80 °C
Temperature (cable mount PRT) -50…100 °C
Accuracy
Frost/Dew point ≤ ± 0.1 °C
Temperature ± 0.07 °C
Reproducibility
Frost/Dew point ≤ ± 0.07 °C
Digital I/O RS-232
Mirror temperature sensor Platinum Resistance Thermometer (Pt100)
Gas connections
Transport case
Calibration: 5 points
Valid calibration certificates from an accredited calibration laboratory or National Metrology Institute shall be provided.
14 | Laptop Computer | Minimum Screen Size: 15 inches | 1 | 1.5
Minimum Processor Speed: 2.6 GhZ
Minimum RAM: 6GB
Minimum Hard Drive: 500 GB
Optical DVD Drive
English Keyboard
Operating System: Windows 7 or higher
15 | Laboratory environmental condition sensors | Temperature: +5 °C to +45°C d=0,01°C accuracy:0,1°C | 1 | 2
Humidity. 20% to 100%d=0,1% accuracy: 2%
With data logger
RS232 connection
A valid calibration certificate from an accredited calibration laboratory or National Metrology Institute shall be provided for the system with inserts.
319.5
Lista de equipos requeridos para el Laboratorio de Volumen
No. | Item | Specifications | Qty. | Budget | Comments:
2 | Balance | d=0,01 mg | 1 | 26 | See: XPE205 by Mettler Toledo
Capacity: 220 g
Repeatability:
3 | Balance | d= 0,1 mg | 1 | 18 | See: XPE504DR by Mettler Toledo
Capacity: 520 g
4 | Balance | d: 1mg | 1 | 12 | See: WP1203 by Mettler Toledo
Capacity: 1210 g
Repeatability: 0,4 mg
5 | Electronic Thermometer | Sensor: Platinum resistance thermometer | 1 | 3.5 | See: Fluke 1523, 1524 Reference Thermometers or equivalent.
d= 0,001 °C
Accuracy: 0,01 °C
Sheath: 300mmx4mm
A valid calibration certificate from an accredited calibration laboratory or National Metrology Institute shall be provided for the system with inserts.
6 | Laboratory environmental condition sensors | Temperature: +5 °C to +45°C d=0,01°C accuracy:0,1°C | 1 | 2 | See: Opus 20 by Lufft Mess
Pressure: 850 hPa to 1050 hPa d=0,1 hPa accuracy: 2hPa
Humidity: 20% to 100%d=0,1% accuracy: 2%
With data logger
RS232 connection
A valid calibration certificate from an accredited calibration laboratory or National Metrology Institute shall be provided for the system with inserts.
7 | Anti-vibratory table | Dimensions of a table SAL/H: 770 mm (length) x 650 mm (width) x 815 mm (height) | 3 | 15 | Model: SAL/H by Radawag
Dimensions of a granite top: 270 mm (length) x 410 mm (width) x 115 mm (height)
Top made of stainless acid
proof steel.
8 | Mass weights | Set weights OIML Class E2, 1mg-100g | 1 | 2 | See: Haffner or equivalent
A valid calibration certificate from an accredited calibration laboratory or National Metrology Institute shall be provided for the system with inserts.
9 | Mass weights | Set weights OIML Class F1 1mg-1kg | 1 | 6 | See: Haffner or equivalent
A valid calibration certificate from an accredited calibration laboratory or National Metrology Institute shall be provided for the system with inserts.
5 | Laptop Computer | Minimum Screen Size: 15 inches | 1 | 1.5
Minimum Processor Speed: 2.6 GhZ
Minimum RAM: 6GB
Minimum Hard Drive: 500 GB
Optical DVD Drive
English Keyboard
Operating System: Windows 7 or higher
10 | Water distillation System | Capacity: 1 L/hr, Output water conductivity: Class 3, Conductivity <0,5 mS/m | 1 | 6
92
Lista de equipos requeridos para el Laboratorio de tiempo y frecuencia
1 | Laboratory environmental condition sensors | Temperature: +5 °C to +45°C d=0,01°C accuracy:0,1°C Pressure: 850 hPa to 1050 hPa d=0,1 hPa accuracy: 2hPa | 1 | 2
Humidity. 20% to 100%d=0,1% accuracy: 2%
With data logger
RS232 connection
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
2 | Laptop Computer | Minimum Screen Size: 15 inches | 1 | 1.5
Minimum Processor Speed: 2.6 GhZ
Minimum RAM: 6GB
Minimum Hard Drive: 500 GB
Optical DVD Drive
English Keyboard
Operating System: Windows 7 or higher
3 | NPT server | Network Protocol Time server | 1 | 5 | See: Meinberg model M600
Selectable Reference Sources:
GPS/ GLN/GLONASS satellite receiver (L1 frequency band
MRS: (GPS, PPS, 10MHz, NTP): Multi Reference Source - several reference sources, adjustable following priority of signal
Synchronization of NTP and SNTP compatible clients
Supported networking protocols: IPv4, IPv6, HTTPS, HTTP, SSH, TELNET, SCP, SFTP, FTP, SYSLOG, SNMP
Alert-Notification system of status change by Email, WinMail, SNMP or an external connected display
Full SNMP v1,v2,v3 support with own SNMP-daemon for status and configuration and SNMP Trap messages
USB Port for installing firmware updates, Meinberg GPS Antenna/Converter Unit connected with up to 300m of standard coaxial cable RG58
5 | Universal Counter | Universal Counter | 1 | 6 | See: Fluke el PM6690 or
Keysight el 53132A
6 | LED Display Clock | Large LED Displays available as NTP slave clock The LE Displays can be configured to operate as a free running stand alone clock, as a reference clock | 1 | 0.7 | See: Meinberg model:DU70 display clock
LED Display for Time and Date
Stand Alone Clock
RS232 interface
20 mA Interface
RoHS compliant
15.2
Lista de equipos requeridos para el Laboratorio magnitudes eléctricas
1 | Laboratory environmental condition sensors | Temperature: +5 °C to +45°C d=0,01°C accuracy:0,1°C Pressure: 850 hPa to 1050 hPa d=0,1 hPa accuracy: 2hPa | 2 | 2
Humidity. 20% to 100%d=0,1% accuracy: 2%
With data logger
RS232 connection
A valid calibration certificate from an accredited calibration laboratory or National Metrology Institute shall be provided for the system with inserts.
2 | Laptop Computer | Minimum Screen Size: 15 inches | 1 | 1.5
Minimum Processor Speed: 2.6 GhZ
Minimum RAM: 6GB
Minimum Hard Drive: 500 GB
Optical DVD Drive
English Keyboard
Operating System: Windows 7 or higher
3 | Reference Multimeter | Voltage DC Range: 0 to ± 1050 V | 1 | 18 | See: fluke 8508A
Voltage AC Range: 2 mV to 1050 V, 1 Hz to 1 MHz
Current DC
Range: 0 to ± 20 A
Current AC
Range: 2 µA to 20 A, 1 Hz to 100 kHz
Resistance Range: 0 to 20 GΩ
Temperature Range: Range: -200 °C to 660 °C
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
4 | 50-turn current coil | Number of Turns: 50 | 0 | 2.5 | See: Fluke 5500A/coil
Maximum Current
11A rms, continuous
20A rms, 2 minutes
Maximum Duty Cycle Derating
< 11A, continuous
> 11A, 2 minutes ON, 8 minutes OFF
Maximum Voltage: 3V rms
Frequency of Operation DC, 45-440 Hz
For the Current Output: DC, 45-65 Hz, 0-11A 65-440 Hz, 0-2.19999A
Minimum Inside Diameter of Clamp Jaws: 2.54 cm
24
Lista de equipos requeridos para el Laboratorio de longitud
1 | Sine bar | Nominal Measurement: 200 mm | 1 | 5
Measurement
Upper Surface (L×W): 300X100 mm
Inclination Angle (degrees): 0° to 90°
Roller Center Distance: 300 mm
Parallelism: 3 µm
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
2 | Set of angular Block | 6 Blocks (in degrees): | 1 | 3 | See Starret catalog number:
|  | 1°, 3°, 5°, 15°, 30°, and 45° |  |  | AG 16.C
5 Blocks (in minutes):
1', 3', 5', 20', and 30'
5 Blocks (in seconds):
1", 3", 5", 20", and 30"
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
3 | Dial gage comparator | Range: 0-25mm | 1 | 3.5 | See Mitutoyo model UDT 2
Clamping stem diameter: 6mm and 8mm
Graduation .001 mm
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
4 | Granite reference surface | LxW:600 mm x 1000 mm | 1 | 2.5 | See Mitutoyo model:
|  | H:80 mm |  |  | 517-843
14
Lista de equipos requeridos para el Laboratorio de Fuerza y Torque
1 | Force and torque measurement indicator |  | 1 | 17 | See: HBM MGC plus
Data acquisition system
With amplification module for force and torque transducers
High resolution of 1,000,000, an accuracy of 0.0025%
Excellent long-term stability.
2 | Tension Transducers | For calibration of testing material | 2 | 9 | See: HBM model STZ
Load: 600 kN and 1 MN
Class 05 according ISO376
Accessories: Hoods
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
3 | Compression transducers | For calibration of concrete testing machines | 2 | 8 | See: HBM model KDB
Load: 1 MN and 2 MN
Class 05 according ISO376
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
4 | Torque wrench calibrator | Loading Wrench | 0 | 9 | See:Monutz LBF250
Torque Range: Max 250 lbf.ft (3000 lbf.in)
Reaction Position: Horizontally 4" to 28" from the center of the turn-table. Vertically 1" to 4" from the turntable base.
Weight 53 lbs.
Size: 37 1/2” W x 12” H x 10 3/4” D
6 | Torque indicator and sensor | Acurracy: 05% FS | 5 | 15 | See Mountz BMX sensors
50- 500 lbf.in 3/8 inch
75 - 750 lbf.in 3/8 inch
5 - 50 lbf.ft 3/8 inch
10 - 100 lbf.ft 1/2 inch
25 - 250 lbf.ft 1/2 inch
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
58
Lista de equipos requeridos para el Laboratorio de Físico-Química
1 | pH reference standards | pH: 4/7/10/12 U: 0,002 | 4 | 2 | See: Fisher/Merk
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
2 | Electrolytic conductimetry standards | 1,3/ 5/ 15/ 84/ 100/ 147/ 1413/12880 µ | 8 | 2 | See: Fisher/Hamilton/Merk
U=0,25%
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
3 | Viscosity standards | Set of Kinematic viscosity standards | 10 | 5 | See: Cannon Instrument
4 | Chronometer | Digital indication | 1 | 0.2
0,01 s
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
5 | Precision Thermometer | Dual PT100 sensors with digital indicator | 1 | 3.5 | See: Dostmann P795 Dual Channel Digital Thermometer by Cannon Instrument
Wide temperature range: −200 °C to +850 °C
Accuracy: ± 0.015 °C from −50 °C to +200 °C
Resolution: 0.001 °C from −200 °C to +200 °C
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
6 | Viscosity thermostatic bath | For Kinematic Viscosity from 20 °C to 100 °C | 1 | 10 | See Cannon Instrument CT-500
ASTM D445/D446, ISO 3104/3105
Up to 12 inch deep
Up to 7 simultaneous viscosity determinations.
Stability 0,01 °C
7 | Holmium glass filter | Primary Usage: Assessment of wavelength scale accuracy in the UV and visible regions. | 1 | 4
Useable range: 270 nm to 640 nm, instruments with SBW of less than 10 nm.
Physical Configuration: Glass filters 'stress free' mounted in anodised aluminium holder.
A valid calibration certificate from an accredited calibration laboratory1 or National Metrology Institute shall be provided for the system with inserts.
8 | Balance | d= 0,1 mg | 1 | 18
Capacity: 520 g
Designed for hydrostatic weighing
44.7
Other Equipment |  |  |  | 11.5

## AP
InterAmerican Development Bank
SU-L1043
Adquisitions Plan (AP)
2014-06-01
id. | Comp. | Description/Type of Contract | Estimated Contract Cost (US) | Source |  | Procurement Method (1) | Revision (2) |  |  |  |  | Estimated Dates | [6 empty] | Status (3) | Comments
|  |  |  | IDB | Local |  |  | TR ou especificações técnicas prontas | Não-objeção do BID (TR/Edital) | Montagem do processo licitatório | Tramitação na PGE | Publication | Abertura da Licitação | Tramitação do processo licitatório | Adjudicação e assinatura do contrato | Emissão de ordem de serviço | PRAZO DE EXECUÇÃO DO CONTRATO | Completion of Contract
1.CONSULTANCY SERVICES
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.1.1. Specialized Consultancies for Accreditation | 50000 | 100% | 0% | IICQ | Ex-Ante | 4177900% | 4178900% | 4180900% | 4183900% | 2017-02-15 | 2017-02-25 | 2017-04-06 | 2017-04-16 | 2017-04-24 | 2 anos | Ago-17
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.1.1. Specialized Consultancies for Accreditation | 50000 | 100% | 0% | IICQ | Ex-Ante | 4177900% | 4178900% | 4180900% | 4183900% | 2017-08-14 | 2017-08-24 | 2017-10-03 | 2017-10-13 | 2017-10-21 | 2 anos | 2018-02-10
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.1.1. Specialized Consultancies for Accreditation | 50000 | 100% | 0% | IICQ | Ex-Ante | 4177900% | 4178900% | 4180900% | 4183900% | 2018-02-10 | 2018-02-20 | 2018-04-01 | 2018-04-11 | 2018-04-19 | 2 anos | 2018-08-09
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.1.1. Specialized Consultancies for Accreditation | 50000 | 100% | 0% | IICQ | Ex-Ante | 4177900% | 4178900% | 4180900% | 4183900% | 2018-08-09 | 2018-08-19 | 2018-09-28 | 2018-10-08 | 2018-10-16 | 2 anos | 2019-02-05
I.A | 1 | IINDIVIDUAL CONSULTANT Activity 1.2.1. Short Training Courses | 10000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-02-15 | 2017-02-25 | 2017-04-06 | 2017-04-16 | 2017-04-24 | 2 anos | 2017-04-16
I.A | 1 | IINDIVIDUAL CONSULTANT Activity 1.2.1. Short Training Courses | 10000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-04-01 | 2017-04-11 | 2017-05-21 | 2017-05-31 | 2017-06-08 | 2 anos | 2017-07-15
I.A | 1 | IINDIVIDUAL CONSULTANT Activity 1.2.1. Short Training Courses | 25000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-05-16 | 2017-05-26 | 2017-07-05 | 2017-07-15 | 2017-07-23 | 2 anos | 2017-10-13
I.A | 1 | IINDIVIDUAL CONSULTANT Activity 1.2.1. Short Training Courses | 25000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-06-30 | 2017-07-10 | 2017-08-19 | 2017-08-29 | 2017-09-06 | 2 anos | 2018-01-11
I.A | 1 | IINDIVIDUAL CONSULTANT Activity 1.3.1. Awarennes and Sensitivization Workshops | 20000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-09-28 | 2017-10-08 | 2017-11-17 | 2017-11-27 | 2017-12-05 | 2 anos | 2017-12-27 |  | #VALUE! | #VALUE! | 2 anos | Ago-17
I.A | 1 | IINDIVIDUAL CONSULTANT Activity 1.3.1. Awarennes and Sensitivization Workshops | 20000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-12-27 | 2018-01-06 | 2018-02-15 | 2018-02-25 | 2018-03-05 | 2 anos | 2018-03-27
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.3.2. Specialized consultancy for information and analysis | 50000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-02-15 | 2017-02-25 | 2017-04-06 | 2017-04-16 | 2017-04-24 | 2 anos | 2017-03-17
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.3.4. Design and Implementation of Website | 25000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-06-01 | 2017-06-11 | 2017-07-21 | 2017-07-31 | 2017-08-08 | 2 anos | 2017-08-30
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.3.5. Operation of Network | 40000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-06-01 | 2017-06-11 | 2017-07-21 | 2017-07-31 | 2017-08-08 | 2 anos | 2017-11-28
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.4.1. Short Training Seminars | 25000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-02-20 | 2017-03-02 | 2017-04-11 | 2017-04-21 | 2017-04-29 | 2 anos | 2017-05-21
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.4.1. Short Training Seminars | 25000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-05-21 | 2017-05-31 | 2017-07-10 | 2017-07-20 | 2017-07-28 | 2 anos | 2017-08-19
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.4.2. Awarennes Campaings | 25000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-08-30 | 2017-09-09 | 2017-10-19 | 2017-10-29 | 2017-11-06 | 2 anos | 2017-11-28
I.A | 1 | INDIVIDUAL CONSULTANT Activity 1.4.3. Awarennes Campaings | 50000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-08-30 | 2017-09-09 | 2017-10-19 | 2017-10-29 | 2017-11-06 | 2 anos | 2017-11-28
I.A | 3 | INDIVIDUAL CONSULTANT Activity 3.1.1. Best Practices on IPAS | 50000 | 100% | 0% | IICQ | Ex-Ante | 2014-05-20 | 2014-05-30 | 2014-06-19 | 2014-07-19 | 2017-05-21 | 2017-05-31 | 2017-07-10 | 2017-07-20 | 2017-07-28 | 2 anos | 2017-08-19
I.C. | 1 | CONSULTANT FIRM. Activity 1.5.1. Specialized consultancy for strategic plan for the GNBS | 50000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-05-21 | 2017-05-31 | 2017-07-10 | 2017-07-20 | 2017-07-28 | 2 anos | 2017-08-19
I.C. | 1 | CONSULTANT FIRM. Activity 1.5.2. Specialized Training | 25000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-05-21 | 2017-05-31 | 2017-07-10 | 2017-07-20 | 2017-07-28 | 2 anos | 2017-08-19
I.C. | 1 | CONSULTANT FIRM. Activity 1.5.2. Specialized Training | 25000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-07-20 | 2017-07-30 | 2017-09-08 | 2017-09-18 | 2017-09-26 | 2 anos | 2017-10-18
I.C. | 2 | CONSULTANT FIRM. Activity 2.1.2. Specialized Consultancy for architectural and Engineering Designs | 200000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-02-20 | 2017-03-02 | 2017-04-11 | 2017-04-21 | 2017-04-29 | 2 anos | 2017-06-20
I.C. | 2 | CONSULTANT FIRM. Activity 2.1.2. Civil Works Investigation and Supervision | 200000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-04-21 | 2017-05-01 | 2017-06-10 | 2017-06-20 | 2017-06-28 | 2 anos | 2017-08-19
I.C. | 3 | CONSULTANT FIRM. Activity 3.1.2. Redesign of Go Invest Trade and Investment Framework | 510000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-04-21 | 2017-05-01 | 2017-06-10 | 2017-06-20 | 2017-06-28 | 2 anos | 2017-08-19
I.C. | 3 | CONSULTANT FIRM. Activity 3.1.4. Website Improvement | 70000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-04-21 | 2017-05-01 | 2017-06-10 | 2017-06-20 | 2017-06-28 | 2 anos | 2017-08-19
I.C. | 3 | CONSULTANT FIRM. Activity 3.2. Export and Investment Strategy | 400000 | 100% | 0% | QCBS | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-05-21 | 2017-05-31 | 2017-07-10 | 2017-07-20 | 2017-07-28 | 2 anos | 2017-09-18
IV.B |  | A. INDIVIDUAL CONSULTANT: Midterm and Final Evaluation | 30000 | 0% | 100% | IICQ | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-07-20 | 2017-07-30 | 2017-09-08 | 2017-09-18 | 2017-09-26 | 2 anos | 2017-10-18
IV.B |  | B. CONSULTANT FIRM: Baseline collection and surveys | 15000 | 0% | 100% | DC (5) | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-07-20 | 2017-07-30 | 2017-09-08 | 2017-09-18 | 2017-09-26 | 2 anos | 2017-10-18
IV.B |  | C. INDIVIDUAL CONSULTANT: Ex-post cost benefit analysis | 20000 | 0% | 100% | NICQ | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-01-15 | 2017-01-25 | 2017-03-06 | 2017-03-16 | 2017-03-24 | 2 anos | 2018-12-15
TOTAL 1 |  |  | 2145000
GOODS AND SERVICES
II.B | 2 | Software Equipment for CUS | 19000 | 100% | 0% | PC | Ex-ante | 2014-10-20 | 2014-10-30 | 2014-11-19 | 2014-12-19 | 2017-11-15 | 2017-11-25 | 2018-01-04 | 2018-01-14 | 2018-01-22 | 2 anos | 2017-12-17
II.B | 2 | Office Equipment | 50000 | 100% | 0% | PC | Ex-ante | 2014-10-20 | 2014-10-30 | 2014-11-19 | 2014-12-19 | 2017-11-15 | 2017-11-25 | 2018-01-04 | 2018-01-14 | 2018-01-22 | 2 anos | 2017-12-17 |  | Two lines
IIIi.A | 2 | Laboratory Equipment | 1544500 | 100% | 0% | ICB | Ex-ante | 2014-10-20 | 2014-10-30 | 2014-11-19 | 2014-12-19 | 2017-02-15 | 2017-02-25 | 2017-04-06 | 2017-04-16 | 2017-04-24 | 2 anos | 2017-12-17
| 3 | 3.1.5. Boutique Forum | 90000 | 100% | 0% | NCB | Ex-ante | 2014-10-20 | 2014-10-30 | 2014-11-19 | 2014-12-19 | 2018-02-15 | 2018-02-25 | 2018-04-06 | 2018-04-16 | 2018-04-24 | 2 anos | 2018-04-16
IIII.B | 3 | 3.1.3. Customer Relationship Management Software | 40000 | 100% | 0% | ICB | Ex-ante | 2014-10-20 | 2014-10-30 | 2014-11-19 | 2014-12-19 | 2017-02-15 | 2017-02-25 | 2017-04-06 | 2017-04-16 | 2017-04-24 | 2 anos | 2017-12-17
IIII. C | 2 | Laboratory Furniture | 100000 | 100% | 0% | ICB | Ex-ante | 2014-10-20 | 2014-10-30 | 2014-11-19 | 2014-12-19 | 2017-02-15 | 2017-02-25 | 2017-04-06 | 2017-04-16 | 2017-04-24 | 2 anos | 2017-12-17
TOTAL 2 |  |  | 1843500
3. CIVIL WORKS
I.A | 1 | 1.3.1.2 Facility for the GNBS | 3500000 | 100% | 0% | ICB | Ex-ante | 2014-05-02 | 2014-05-12 | 2014-05-17 | 2014-06-16 | 2017-01-15 | 2017-01-25 | 2017-03-06 | 2017-03-16 | 2017-03-24 | 2 anos | 2017-12-16
TOTAL 3 |  |  | 3500000
SUBTOTAL ACQUISITION PLAN |  |  | 7488500
1 | Goods and Works: ICB: International competitive bidding; LIB: limited international bidding; NCB: national competitive bidding; PC: price comparison; DC: direct contracting; FA: force account; PSA: Procurement through Specialized Agencies; PA: Procurement Agents; IA: Inspection Agents; PLFI: Procurement in Loans to Financial Intermediaries; BOO/BOT/BOOT: Build, Own, Operate/Build, Operate, Transfer/Build, Own, Operate, Transfer; PBP: Performance-Based Procurement; PLGB: Procurement under Loans Guaranteed by the Bank; PCP: Community participation procurement. Consulting Firms: QCBS: Quality- and Cost-Based Selection QBS: Quality-Based Selection FBS: Selection under a Fixed Budget; LCS: Least-Cost Selection; CQS: Selection based on the Consultants’ Qualifications; SSS: Single-Source Selection. Individual Consultants: NICQ: National Individual Consultant selection based on Qualifications; IICQ: International Individual Consultant selection based on Qualifications
2 | The Bank's new default methodology is ex post, based on risk assessment.
3 | This column “Status” should be used for retroactive procurement and for procurement plan updates.
7669
7790500
Formulas:
Abertura da Licitação = [Estimated Dates]+10
Tramitação do processo licitatório = [Abertura da Licitação]+40
Adjudicação e assinatura do contrato = [Tramitação do processo licitatório]+10
Emissão de ordem de serviço = [Adjudicação e assinatura do contrato]+8
Estimated Dates = +[Estimated Dates, row above]+180
Estimated Dates = +[Estimated Dates, row above]
Completion of Contract = +[Estimated Dates]+180
Completion of Contract = +[Completion of Contract, row above]+90
Não-objeção do BID (TR/Edital) = [TR ou especificações técnicas prontas]+10
Montagem do processo licitatório = [Não-objeção do BID (TR/Edital)]+20
Tramitação na PGE = [Montagem do processo licitatório]+30